Table of Contents

1.0	Introduction	1

1.1	Objective	1

1.2	Scope	1

1.3	Background	1

1.4	Definitions, Acronyms, and Abbreviations	2

1.5	References	2

2.0	Procedure	3

2.1	Requirements To Be Tested	3

2.2	Test Methodology	3

2.3	Test Article	4

2.4	Sample Justification	4

2.5	Detailed Test Procedure	4

3.0	Acceptance Criteria	5

4.0	Attachments	6

# Introduction

## Objective

The purpose of this document is to define the design validation testing of the updated automated scoring algorithm using pressure readings known to be “suspect” (physiologically invalid) but not marked “suspect” by the previous scoring algorithm.

## Scope

The scope of this test plan is limited to the updated algorithm that assesses the pulmonary artery (PA) pressure readings in the CardioMEMS HF System, defined in SO-500119, using pressure readings known to be suspect because summary reading values are not consistent with valid pressure measurement values.

## Background

The CardioMEMS HF System includes a scoring algorithm, defined in SO-500119, that identifies suspect readings that are likely to be clinically or physiologically invalid.  If a reading scores above the designated threshold, the pressure reading is flagged and displayed as “suspect” on Merlin.net.  The clinician user can review “suspect” readings and determine if they are clinically valid or not.

CAPA 17056 was initiated to address two complaints related to the scoring algorithm regarding summary reading values that did not follow necessary rules for pulmonary artery pressures.  In one, the systolic, diastolic and mean summary reading values were self-inconsistent, such that they did not follow the rule that systolic ≥ mean ≥ diastolic.  In the second complaint, the summary reading values were not bounded by the waveform, such that a systolic, mean, or diastolic pressure value was larger or smaller than the maximum or minimum point of the pressure waveform.

As detailed in CAPA Measure Phase Memo 90277819, in response to these two complaints, on 04 May 2017, 1,395,000 readings from were analyzed, and 383 readings from 191 sensors were found to have incorrect summary reading values due to one of these two issues (self-inconsistent or outside the bounds of the waveform).

Updates made to the scoring algorithm via CAPA 17056 will be revalidated via design validation protocol 90314713 using set of randomly-sampled, independent pressure readings sourced from the Merlin.net database, which are a mix of valid and suspect readings, with endpoints evaluating Sensitivity and Specificity of the algorithm.  This second validation protocol will apply the updated waveform scoring algorithm to pressure readings identified to have summary pressure values that are self-inconsistent or outside the bounds of the waveform.

## Definitions, Acronyms, and Abbreviations

| Term | Definition |
| --- | --- |
| CRM | Cardiac Rhythm Management |
| Gluon | Open-source deep learning platform on which developers can prototype, build, train, and deploy machine learning models |
| HF | Heart failure |
| HES | Hospital Electronics System |
| PA | Pulmonary artery |
| PES | Patient Electronics System |
| RCO | Remote Care Operation |
| Summary values | The maximum, minimum, and average pressure values calculated from the PA pressure waveform |

Table 1: List of terms defined for use in the following report.

## References

| Number | Title |
| --- | --- |
| 90277819 | CAPA Measure Phase Document |
| 90314713 | CardioMEMS Waveform Scoring Algorithm Design Validation Protocol |
| 90327211 | SW Development Tool: Web App Scoring Validation |
| 90327231 | SW Development Tool: Interrogator Scoring Validation |
| SO-500119 | HF Reading Scoring Algorithm |
| SOP-000136 | Software Tool Control Process |
| TR-1002-101 | DIOVV – HF Electronics |
| TR-1011-008 | i3 DIOVV |

Table 2: List of documents referenced in this test protocol.

# Procedure

## Requirements To Be Tested

| Document | User Need ID | User Need |
| --- | --- | --- |
| TR-1002-101 | 3.A.iv.2 | The system must be capable of displaying pulmonary artery pressure with a high degree of accuracy which shall allow the user to identify clinically significant changes in blood pressure of the heart. |
| TR-1011-008 | 1.A.ii.7 | Pressure measurements shall be recorded as long as the frequency, pressure waveform and signal strength are adequate. |

Table 3: List of user needs relevant to this validation test.

## Test Methodology

The basic steps of the validation are:

- Identify pressure readings for CardioMEMS patients from Merlin.net with summary pressure values that are inconsistent or outside the bounds of the waveform.

- Run readings through the updated software scoring algorithm.

- Verify that all readings are scored as suspect.

Sample Pressure Readings from Merlin.net

The Remote Care Operations group was responsible for providing data from Merlin.net and determining who could access that information.  The RCO query process follows detailed work instruction 60067208, available in the CRM document control system, which sets requirements for management of protected data.  The data provided includes only Sensor ID and Reading ID identifiers, which are necessary to align and adjudicate each reading.  Controlled access is required to associate the sensor and reading ID with additional patient data.

Pressure readings with summary pressure values that are subject to summary reading values that are inconsistent or outside the bounds of the waveform are identified from Merlin.net.

The pressure data for these pressure readings from Merlin.net are exported into a spreadsheet.  That set of pressure data includes numeric values for various characteristics of the pressure waveform, which are already calculated by Electronics and Server scoring algorithm and is unchanged in the updated scoring algorithm.  A library of the associated XML waveforms is also aggregated for all readings in each spreadsheet.

Run Readings through Scoring Algorithm

Scripts will be used to read the data from the waveform library as inputs for the Server and Electronics code that contains the scoring algorithm.  The scripts will also output the results of the scoring algorithm into a CSV file.

Verify Readings are Suspect

The CSV files and scoring algorithm results will be imported back into the spreadsheet.  All pressure readings used for this validation are known to be invalid, and should be scored as “suspect”.

## Test Article

This procedure validates the updated scoring algorithm, defined in SO-500119 Revision G, implemented in the following articles.  See Table 4.

| Test Article | Model Number | Software Model Number | Software Revision |
| --- | --- | --- | --- |
| Backend Server | MN6000 | SO-550095 | L (pre-release) |
| i2 Hospital Electronics System | CM3000 | SO-550142 | D (pre-release) |
| i3 Patient Electronics System | CM1100 | SO-550150 | F (pre-release) |

Table 4: List of equipment evaluated during testing.

## Sample Justification

The number of readings used for this validation test is not statistically based.  All readings from Merlin.net that have either the following characteristics will be used for algorithm validation.

- NOT (systolic ≥ mean ≥ diastolic)

- Systolic, mean, or diastolic values outside the minimum and maximum point of the pressure waveform

## Detailed Test Procedure

- Identify raw pressure data readings.

- Run a script, detailed in Attachment A, on the open source database PostgreSQL to query (pull) data from the CardioMEMS Production Merlin.net reporting database.

  - The script will create an empty template table to be populated with readings from the CMEMS Production Merlin.net reporting database.

  - Readings flagged for presenting incorrect or inconsistent summary values as described in CAPA Measure Phase Memo 90277819 will be accumulated into the template table.  The script will search through readings for all time up to the present.

  - The final outputs will be

    - 2 Excel files of attribute data — one for readings presenting inconsistent summary values and one for readings with incorrectly bounded summary values. A row in each sheet will pertain to a single reading.

    - 2 libraries — one for each Excel file — of XML waveforms associated with the readings in each datasheet

    - 2 manifests — one for each XML library — listing the file path to each XML waveform

- Use the Ruby script described in 90327211 – SW Development Tool: Web App Scoring Validation to run the scoring algorithm and output results within software model SO-550095 (used on the Backend Server). This software tool must be validated according to guidelines for non-product software validation outlined in SOP-000136.

- The script will require the following inputs to run:

  - The two sets of attribute data exported from the data query output file in CSV format

  - The library of associated XML waveform files

    - The included manifest.txt file capturing filenames of all XML waveforms in the library and their associated identifiers

  - The function “reading_xml_parser.rb” to parse the XML library and link each waveform to its attributes in the adjudicated data CSV file

  - The following functions, sourced from the same code base used in product software SO-550095

    - “models.rb” to assign each reading summary values (e.g. systolic pressure, heart rate) and calculate pressure baselines

    - “reading_score.rb,” which contains the algorithm detailed in SO-500119

- The script will parse the waveform XML library for a single file and match its unique Reading ID to the corresponding input CSV row.

- The script will calculate summary and baseline values.  A reading will be ruled “suspect” if its attributes score is equal to or greater than the validity threshold (defined in SO-500119) or “approved” if it scores under the validity threshold.

- The final output for each dataset will be a CSV file showing reading ID, sensor SN, and the numeric algorithm score.

- Copy the algorithm scores from the server CSV output files into the “Server” sheet of the endpoints calculator, Attachment B.

  - The Excel sheet will tally the algorithm output scores and show the proportion of readings correctly identified as “suspect.”

- Run the Gluon script described in 90327231 – SW Development Tool: Interrogator Scoring Validation in Linux to input data and run scoring algorithm on software models SO-550142 (I2 HES) and SO-550150 (I3 PES).  This software development tool must be validated according to guidelines for non-product software validation outlined in SOP-000136.

- The Gluon script will require the following files to run:

  - The set of adjudicated data exported from the “Final” sheet of the data query output file in CSV format

  - The library of XML waveform files

    - The included manifest text file to extract the XML waveform for each reading in the adjudicated data CSV file and run the algorithm

- Open the manifest in Notepad++. Use the Find and Replace function to remove all quotations (“) from the file content.

- In Gluon, enter ./scoringvalidation name.csv to navigate to the adjudicated data input file, where name is the title of the target file

- The script will ask if the manifest text file is stored within the XML library folder. Check and confirm when prompted.

- Enter the file path to the XML library when prompted, following format in the example: /home/tstark/Desktop/

- Enter the file path to the desired output file location when prompted, following format in the example: /home/tstark/Desktop/

  - The script will ask if a file with the intended output title already exists at the specified location. Check and negate when prompted.

- Type the file path to the manifest text file, then enter the file path to the XML library on the next line. Hit enter.

  - The script will navigate to a location specified by the manifest text file and pull a unique Reading ID from the XML file.

  - The script will then access the adjudicated dataset CSV to find the entry with a matching Reading ID.

  - Next, “manifest.txt” will assign the reading summary values and calculate its pressure baselines. The reading will be scored “suspect” if it scores over a threshold value and “approved” if it scores under this threshold value.

  - “manifest.txt” will append the algorithm score to its respective row in the set of input CSV data and output as a new CSV file after all readings have been scored.

- Copy the numeric algorithm scores from the interrogator script CSV output files into the “Interrogator” sheet of the endpoints calculator, Attachment B.

  - The Excel sheet will tally the algorithm output scores and show the proportion of readings correctly identified as “suspect.”

# Acceptance Criteria

Validation will be considered successful if all readings known to be invalid are scored “suspect” in both the Server and Interrogator software.

# Attachments

| Attachment | Title | Description | Program |
| --- | --- | --- | --- |
| A | query.sql | Data query script | PostgreSQL |
| B | Suscheck.xlsx | Calculator for proportion of suspect readings | Excel |

Table 5: List of attachments relevant to this protocol.