| Written By: | Erik Moore |
| --- | --- |

| Software Name: | CardioMEMS HF System Web Apps I3 Patient Electronics System I2 Hospital Electronics System I2 Patient Electronics System |
| --- | --- |

# Introduction

## Purpose

The purpose of this document is to describe the automated algorithm that is used to determine if a particular reading is influenced by an artifact. As readings are processed, any technically invalid characteristics add to the score value of the reading.  When the score value crosses the threshold then the reading is marked for the user to review.  This scoring algorithm is not a replacement for clinical judgment or interpretation of physiological events.  It is designed to assist the physicians in their review of clinical data by identifying readings that may not be physiologic.

## Scope

Within the CardioMEMS HF system, the scoring algorithms are used to identify and screen for readings that are technically influenced by artifacts, patient movements, etc.  The output of the scoring is simply the determination of whether or not the reading is free from technical faults, e.g., artifacts, patient movements, etc.  The scoring algorithm does not generate any physiological parameters.

The same scoring algorithm described below is used in multiple places in the CardioMEMS HF system.  It runs on the Hospital Electronics System (HES) and the Patient Electronics System (PES).  This allows the user to immediately respond to a suspect reading by capturing an additional reading.  The algorithm also runs on the Web Application.  This allows the system to take into account user approved/rejected readings in the multi-day averaged scoring components.

On the HES that is used in the hospital by clinicians, the software uses the scoring algorithm to prompt the clinician to take another reading if a reading is determined to be technically invalid.  Invalid readings are not transmitted to the website.

On the PES that is used in the home by the patient, the software uses the scoring algorithm to prompt the patient to take a second reading if the first reading is determined to be technically invalid.  Whether or not the second reading is determined to be technically valid, both the technically invalid reading and the second reading are transmitted to the database server and the patient is not prompted to take another reading.

On the Web Application, the same scoring algorithm is used to identify readings with technical faults.  If the reading does not meet the criteria, then the reading is flagged as suspect and an indication is displayed on the trend graph that a suspect reading has been received and a facility to allow site review of said reading is presented.

Both the mean and pulse trends rely on trends made from averaging accepted readings.  On the Web Application, the system accepts or rejects the reading but the user can modify this decision.  When calculating mean and pulse trends, the PES synchronizes the state of its readings with the Web Application to exclude user-rejected readings.

## Definitions, Acronyms and Abbreviations

| Term | Definition |
| --- | --- |
| AvgSS | Average signal strength in reading |
| CP | Number of consecutive pressure data points |
| HES | Hospital Electronics System |
| HR | Heart Rate |
| Max waveform | Maximum pressure value in the waveform, rounded up to the nearest integer |
| Min waveform | Minimum pressure value in the waveform, rounded down to the nearest integer |
| Min() | Results in the minimum value of the two listed. |
| MSS | Minimum signal strength in reading |
| P | Pressure data point |
| PES | Patient Electronics System |
| PP | Pulse pressure |
| Trend | A seven day running average of the given parameter taken from accepted readings. (NOTE that the Patient Electronics System will wait 7 days from the first reading to use the trend. This allows the trend to stabilize.) |

## Document Overview

Section 2.0 is an overview of the components of the scoring algorithm.

Section 3.0 provides the details regarding the scoring algorithm.

## Background

The US Feasibility portion of the CHAMPION clinical trial for the CardioMEMS HF system started in December 2006.  Initially, during the feasibility and pivotal clinical trial, someone from the company reviewed each waveform visually to determine if it was technically valid.  The procedures used for these assessments were CardioMEMS Home Monitoring System Database (SOP-000049, now obsolete/ historical) and Technical Validity of Home Monitoring Database Readings (SOP-000055, now obsolete/ historical).  After the review any technically invalid readings were marked electronically.

In anticipation of the commercial release of the CardioMEMS HF System, a software algorithm was developed to perform the technical validity check automatically.  The automatic algorithm was optimized against the electronic manual results and revalidated.

In 2017, to address improvements from CAPA 115297, checks were added to check the maximum and minimum values of the waveform against the systolic, diastolic and mean pressure components.  Scoring was added to the hospital electronics system.  Validation was repeated.

Original version:

- 06-Apr-2009 – ECN-01508, Version A of SO-500119

- 29-May-2009 – ECN-01601, Version A of validation protocol and results SO-500132/ SO-500133

Optimized version:

- 19-Oct-2010 – ECN-02135, Version C of SO-500119

- 19-Oct-2010 – ECN-02135, Version B of validation protocol and results SO-500132/ SO-500133

2017 version

- 11-Dec-2017 – C326978, Version F of SO-500119

- 24-Apr-2018 – Validation Protocol 90314713/ Report 90356421 and Validation for Known Suspect Readings Protocol 90332959/ Report 90356423 (all Version A)

Historical notes regarding other versions of SO-500119:

- Version B was the same version of the algorithm as Version A.  Only descriptive text was changed.

- Version D updated the scoring on the Electronics System to be the same as the version running on the CardioMEMS HF System Web Apps.

- Version E was a text clarification update that also fixed a typo in the diastolic pressure component.

- Version G clarified that the scoring only failed if sum of components was ≥ 40.

- Version H explained that the patient electronics system downloads the recent reading state from the website, added to the background, and explained that the hospital electronics system does not use mean and pulse difference components.

# Scoring Component Overview

## Scoring Components

The score given to the reading is made up of the following components.

- Sanity Check – Ensures that the summary properties of the reading are in physiological ranges.

- Minimum Signal Strength – Checks to ensure the minimum signal strength is in an acceptable range.

- Average Signal Strength – Verifies that the average signal strength is in an acceptable range.

- Pulse Difference – Inspects the pulse pressure for large deviations from the 7 day moving average of accepted readings.

- Mean Difference – Registers if the mean pressure deviates from the 7 day moving average of accepted readings by a large margin.

- Pressure Spike – Examines for any pressure deviations outside of physiological range between consecutive pressure data points.

- Mean Pressure Plateau –The pressure waveform is checked for features that are outside a normal range for an extended period of time.

Total Score:  The total score is the sum of all of the above score factors.  If the total score is equal to or greater than the total threshold listed below in the table, then the reading will be flagged.

## Scoring Components Summary

Table 1: Scoring Components

| Component | Criteria | HF Software Score |
| --- | --- | --- |
| Sanity Check |  |  |
| 1. Heart Rate | > 150 or ≤ 0 | 60 |
| 2. Mean Pressure | > 150 or ≤ 0.1 or < diastolic or > systolic or > max waveform or < min waveform | 60 |
| 3. Diastolic Pressure | > 100 or ≤ -5.0 or > max waveform or < min waveform | 60 |
| 4. Systolic Pressure | >200 or ≤ 0.1 or > max waveform or < min waveform | 60 |
| Minimum Signal Strength |  |  |
|  | < 60 | Min(60 – MSS, 20) |
| Average Signal Strength |  |  |
|  | < 80 | Min(80 – AvgSS, 20) |
| Pulse Difference |  |  |
|  | > trend * 3.0 | 7 * (PP – (trend * 3.0)) |
|  | < trend * 0.5 | 7 * (trend – PP) |
| Mean Difference |  |  |
|  | > trend * 1.8 | 4* (mean – (trend * 1.8)) |
| Pressure Spike |  |  |
|  | n,n+4> 40 | 1.8 * n,n+4- 40 |
| Mean Deviation Plateau |  |  |
|  | P > ½ PP + mean or P < mean – ½ PP for > (2 * HR) number of consecutive points | CP / 6 |
| Total Threshold |  |  |
|  | Sum of components < 40 | Valid |
|  | Sum of components ≥ 40 | Invalid |

*Note: Max waveform is rounded up to the nearest integer.  Min waveform is rounded down to the nearest integer.  Systolic, diastolic, mean and heart rate values are integers.

*Note: The mean and pulse difference components are not used in the hospital electronics system, since these trends do not exist at implant which is when the hospital electronics system is principally used.

# Scoring Component Details

## Sanity Check

The sanity check is made up of several sub-components.

- Heart Rate – a heart rate outside of the acceptable range is given a scoring component above the total threshold; otherwise, the scoring component is zero.

- Mean Pressure – a mean pressure outside the acceptable range or the range diastolic pressure – systolic pressure or outside of the bounds of the waveform is given the scoring component above the total threshold; otherwise, the scoring component is zero.

- Diastolic Pressure – a diastolic pressure outside the acceptable range or outside of the bounds of the waveform is given a scoring component above the total threshold; otherwise, the scoring component is zero.

- Systolic Pressure – a systolic pressure outside the acceptable range or outside of the bounds of the waveform is given a scoring component above the total threshold; otherwise, the scoring component is zero.

## Minimum Signal Strength

If the signal strength minimum for the reading is greater than or equal to the criteria in Table 1, a scoring component of zero is given; otherwise, the scoring component is determined to be the value of the scoring component referenced in Table 1.

## Average Signal Strength

If the average signal strength for the reading is greater than or equal to the criteria in Table 1, a scoring component of zero is given; otherwise, the scoring component is the value of the scoring component referenced in Table 1.

## Pulse Pressure Difference

A trend of pulse pressures for the patient is determined by averaging accepted readings taken in the last seven days. If the pulse pressure exceeds the criteria listed in Table 1, the scoring component is the value of the scoring component referenced in Table 1. Otherwise the scoring component is zero.

## Mean Pressure Difference

A trend of mean pressures for the patient is determined by averaging accepted readings taken in the last seven days. If the mean pressure exceeds the criteria listed in Table 1, the scoring component is the value of the scoring component referenced in Table 1.  Otherwise the scoring component is zero.

## Pressure Spike

The pressure data is analyzed and the largest delta between two points separated by a distance of four points is determined. If the pressure spike exceeds the criteria listed in Table 1, the scoring component referenced in Table 1 is added to the score.  Otherwise the scoring component is zero.

## Mean Deviation Plateau

The pressure data is analyzed to find the longest number of consecutive data points that exist outside of a range which is defined to be mean pressure - ½ pulse pressure to ½ pulse pressure + mean pressure. If the mean deviation plateau exceeds the criteria listed in Table 1, the scoring component referenced in Table 1 is added to the score.  Otherwise the scoring component is zero.